[image: image1.png]



[image: image2.png]



NOAA G-IV
WINTER STORMS 2006
Flight 3 – Track 54
PHNL – PHNL
Flight ID: 060128N
Sensor or system





Number or Name
Inertial (for wind derivation)  



VEW_PITR/VNS_PITR


Temperature Probe





AT3
Dew Point Probe





DPLC
Static and Dynamic Pressure




PS2C/QC2C
Project Directory / Constants File



601 / 49cal051
Notes:

There were no time/data glitches during this flight.  However, there was a gap/spike in the GPS data which began at 2122Z.  Turbulence was encounter around 0026Z-0029Z.  This was reflected in the data.
AT2 was not installed for these missions.  DPLC was chosen as the primary dewpoint sensor because DPRC was greater than the ambient temperature during ascent and descent.  It also read 5-10 degrees warmer than DPLC through most of the flight.  Because of the ascent data, the data from DPLC appears to be the more accurate of the 2 sensors. 

During ascent and descent, dynamic and differential pressures diverged due to location and air flow around the aircraft.  All data, specifically pressure and wind data, should be used carefully during these times.  
After 2000Z, DAP1 was no longer operational.  Raw dynamic attack pressure right (DAP2) was and should be used for derivations.
All other instruments worked properly.





Takeoff


Landing

Aircraft Static Pressure (mb)

1019.1 mb

1016.8mb

Station Altimeter


30.09”


30.02”

Flight Director:

Contact Paul Flaherty
Phone #:


(813) 828-3310 ext.  3094






