| Cal Update | Description | Manufacturer | Model | Serial Number | C0 Slope | C1 Offset |
| --- | --- | --- | --- | --- | --- | --- |
|  | Flow Angle (DAP) | Rosemount | 1221F2AF7B1B | 842 | 17.2537 | - 0.1021 |
|  | Flow Angle (DBP) | Rosemount | 1221F2AF7B1B | 861 | 17.1471 | 0.1028 |
|  | Flow Angle (AP) | Rosemount | 1221F2VL7B1B | 877 | 6.8881 | 0.0641 |
|  | Flow Angle (BP) | Rosemount | 1221F2VL7B1B | 878 | 6.8910 | - 0.0437 |
|  | Copilot (PQ3/PQF2) | Rosemount | 1221F2AF7B1B | 2443 | 17.2312 | 0.8782 |
|  | Copilot (PQ2/PQF1) | Rosemount | 1281B2BEP3 | 161 | 16.5540 | - 0.3379 |
|  | Copilot (PS2/PSF) | Rosemount | 1281B2BEP3 | 161 | 79.9650 | 249.2979 |
|  | Wingtip (PQ) | Rosemount | 1281B2BEP3 | 163 | 16.7537 | - 10.2298 |
|  | Wingtip (PS) | Rosemount | 1281B2BEP3 | 163 | 79.9704 | 249.4970 |
|  | Cabin Pressure | Vaisala | PTB 220 | W3120001 | 120.000 | 500.000 |
|  | Radome Flow Angle (Attack) | Rosemount | 1221F2 VL7B1B | 2317 | 6.8765 | - 0.3298 |
|  | Radome Flow Angle (Sideslip) | Rosemount | 1221F1VL1B1B | 879 | 6.8833 | - 0.0814 |
|  | Radome Flow Angle (Dynamic) | Rosemount | 1221F1AF1B1B | 286 | 34.5926 | - 0.7240 |
|  | Radome Flow Angle (Total) | Rosemount | 1201F2A14A1A | 260 | 206.8766 | - 2.5857 |
|  | Total Temp 1 (TT1) (Fwd) | Rosemount | 102CH2AF w/ 510DY1 | 58171 w/ amp #16 | 12.0289 | - 0.1002 |
|  | Total Temp 2 (TT2) (Aft) | Rosemount | 102CP2AF w/ 510GB41E | A7775 w/ amp #152 | 6.9934 | 0.0565 |
|  | Total Temp 3 (TT3) (Radome) | Rosemount | 102CL2AZ w/ 510GB43E | N/A |  |  |

| Cal Update | Description | Manufacturer | Model | Serial Number | C3 (X3) | C2 (X2) | C1 (X) | C0 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  | CO2 Radiometer | Barnes | PRT-5 | N/A |  |  |  |  |
|  | SST Down Radiometer | Barnes | PRT-5 | 588 | 0.0906 | -0.3626 | 8.1467 | 5.4626 |
|  | SST Up Radiometer | Barnes | PRT-5 | N/A |  |  |  |  |
|  | Dewpoint #1 | General Eastern | 1201 w/ 1011B | 52724 w/ 52722 |  |  | 11.1886 | - 16.8561 |
|  | Dewpoint #2 | EdgeTech | Vigilant w/ 137 | 29833 w/ 1744 |  |  | 20.0 | -60.0 |
|  | Dewpoint/ Water Vapor | MayComm | TDL | 1 |  |  | 7.5 | -25.0 |
|  | Water Vapor | SpectraSensors | WVSS II | 1 |  |  | 8000.0 | 0 |