1

NOAA G-IV ERROR SUMMARY

WINTER STORMS 2007

TRACK56

Flight ID: N070214

Sensor or system						Number or Name

INE	(for wind derivation)				VEW_PITR, VNS_PITR

Accelerometer						ACINS_PITR

Temperature Probe				          AT3

Dew Point Probe						DPRC

Static Pressure						PS2M

Dynamic Pressure						QC2M

Attack Angle						AKRD1

Slip Angle						          SSRD1

Vert. Wind						          PALT

Constants File                                                           49cal061

Project Directory						/proj/701

Local surface observation not current or unavailable.

Notes:

There waere no data gaps.

The Total Temperature #2 (TT2) probe was not available for this mission.

Static Pressure #1...PS1M...exhibited several stair-step jumps at 205138Z, 210721Z,  and 231913Z. This did not impact any meteorological derivations since it was not selected.

All other instruments worked optimally during the flight.

All 16 tempdrop messages were successfully sent to NCEP and made it into the 15/00Z model run.

Takeoff		Landing

Aircraft Static Pressure                1022.2mb	          1019.2mb

Corrected Tower Pressure           1022.3mb	          1019.1mb

Flight Director:	A. Barry Damiano       (813) 828-3310 ext. 3073