NOAA G-IV N49RF ERROR SUMMARY

1

WINTER STORMS 2011

13 Jan 2011 Modified TRACK27 (04WSC)

NCEP (PWT) PANC PANC

Flight ID: 20110113N1

Reference Sensor or System	   			Number or Name

Accelerometer						AccZI.1

Altitude							AltGPS.3

Attack Angle						AA.2

Dew Point Probe						TD.2

Dynamic Pressure						PQF.2

Geopotential Altitude					AltGPS.3

Inertial Selected						INE2

Static Pressure						PSF.2

Slip Angle							SA.1

Temperature Probe				                     TTM.3

True Airspeed						TASF.3

Constants File                                                           49cal102

Flight Directory						acdata/2011/MET/20110113N1

Local Met Data			Takeoff (0901z)		Landing (1632z)

Aircraft Static Pressure		  1018.2 mb			   1015.5 mb

Tower Pressure (corrected)	  1019.0 mb			   1016.2 mb

__________________________________________________________________

Notes:

There were a few data gaps in AltRa.1 from the APN-232 radar altimeter:

09:16:10z

09:17:54-09:18:39z (multiple gaps in this interval)

14:51:33-14:51:34z

15:57:22- 15:57:30z

The ALTPA.1 (NACA Pressure Altitude) runs about 60 meters lower than a tight clustering of Air Data Computer and ADC Baro Corrected sources.

Angle of Attack sources: AaADDU.1 (from the Air Data Computer) was consistently about 0.5 degrees higher than either AA.1 or AA.2, or AaADDU.2. Spike in AA.1 prior to takeoff from 9:00:28 - 9:01:01z. Similar AA.1 spike after landing from 16:32:58 - 16:33:39 (with the sharpest portion of the anomaly at 16:33:10). In addition to the spikes....there was a strange rise and fall of AA.1 earlier on the taxi (from 8:56:08 - 8:57:38) that was too protracted to be called a spike....just anomalous.

GDIFF check: mean of 4677 meters of drift for INE #1 (with respect to the Novatel)

1414 meters for INE#2. In creation of higher MET parameters, INS #2 was selected.

The Left Dew Point sensor (TD.1) again read far too high through the entire flight (well above ambient temperature). TD.2 did a reasonably good job (remaining below ambient temperature) but is also too high during the cruise portion of flight at or above 41,000 feet (where sondes suggest an RH near 5% and the dew points generated by TD.2 yield an RH around 25%).

Vertical Winds look fine during the cruise portion of flight (with a mean of -0.04) but then, as expected, show a significant high bias during descent prior to landing.

All other flight level instruments worked optimally during the flight.

- 19 drop points assigned by NCEP

- 19 AVAPS I dropsondes deployed during the mission

- Of those 19 drops, all 19 were good and transmitted.

- Of the 19 sondes, 17 coded surface winds with the other two coding 12 meters and 14 meters for a LAST WIND remark.

Flight Directors: Richard Henning and Jessica Williams (813) 828-3310 ext. 3086