NOAA G-IV N49RF ERROR SUMMARY

1

WINTER STORMS 2011

20 Jan 2011 TRACK76 (07WSW)

NCEP (PWT) RJTY RJTY

Flight ID: 20110120N1

Sensor or system						Number or Name

Accelerometer						AccZI.1

Altitude							AltGPS.3

Attack Angle						AA.2

Dew Point Probe						TD.2

Dynamic Pressure						PQF.2

Geopotential Altitude					AltGPS.3

Inertial Selected						INE2

Static Pressure						PSF.2

Slip Angle							SA.1

Temperature Probe				                     TTM.3

True Airspeed						TASF.3

Constants File                                                           49cal102

Flight Directory						acdata/2011/MET/20110120N1

Local Met Data			Takeoff (0730z)		Landing (1443z)

Aircraft Static Pressure		  992.8 mb			   997.9 mb

Tower Pressure (corrected)	  994.4 mb			   998.3 mb

__________________________________________________________________

Notes:

The only data gaps were in AltRa.1 from the APN-232 radar altimeter:

7:45:27- 7:46:02z (multiple gaps during this interval)

14:20:21 – 14:20:22z

The ALTPA.1 (NACA Pressure Altitude) ran about 75 meters lower than a tight clustering of Air Data Computer and ADC Baro Corrected sources during cruise.

Angle of Attack sources: AaADDU.1 (from the Air Data Computer) was consistently about 0.5 degrees higher than either AA.1 or AA.2, or AaADDU.2. Spike in AA.1 on takeoff roll at 7:29:41z up to 156.6 degrees 21 seconds before liftoff. Similar AA.1 spike on landing rollout up to 109.7 degrees at 14:43:53

GDIFF check: Mean of 1908 meters of drift for INE #1 (with respect to the Novatel)

1189 meters for INE#2. In creation of higher MET parameters, INE #2 was selected.

The Left Dew Point sensor (TD.1) again read far too high through the entire flight (well above ambient temperature). TD.2 did a reasonably good job but is also too high during the cruise portion of flight at or above 41,000 feet (where sondes suggest an RH near 5% and the dew points generated by TD.2 yield an RH around 25%).

Vertical Winds showed a larger-than-normal low bias during the cruise portion of flight (with a mean of -0.23 m/s) and showed a more significant high bias during descent prior to landing.

All other flight level instruments worked optimally during the flight.

- 12 drop points assigned by NCEP

- 12 AVAPS I dropsondes deployed

- Of those 12 drops, all 12 were good enough to create a WMO message and were transmitted. However, one of the good drops was transmitted with no heights: Drop Pt 2 at 09:06:07z failed at 971 mb.

- Of the 12 good sondes, 11 coded surface winds.

Flight Directors: Richard Henning and Jessica Williams (813) 828-3310 ext. 3086