NOAA G-IV N49RF ERROR SUMMARY

1

WINTER STORMS 2011

21 Jan 2011 Modified TRACK77 (08WSW)

NCEP (PWT) RJTY RJTY

Flight ID: 20110121N1

Sensor or system						Number or Name

Accelerometer						AccZI.1

Altitude							AltGPS.3

Attack Angle						AA.2

Dew Point Probe						TD.2

Dynamic Pressure						PQF.2

Geopotential Altitude					AltGPS.3

Inertial Selected						INE1

Static Pressure						PSF.2

Slip Angle							SA.1

Temperature Probe				                     TTM.3

True Airspeed						TASF.3

Constants File                                                           49cal102

Flight Directory						acdata/2011/MET/20110121N1

Local Met Data			Takeoff (0732z)		Landing (1520z)

Aircraft Static Pressure		  997.8 mb			     999.7 mb

Tower Pressure (corrected)	  999.7 mb			   1000.4 mb

__________________________________________________________________

Notes:

The only data gaps were in AltRa.1 from the APN-232 radar altimeter:

7:49:30-7:50:29z (multiple gaps during this interval)

14:57:47z-14:57:48z

The ALTPA.1 (NACA Pressure Altitude) ran about 65 meters lower than a tight clustering of Air Data Computer and ADC Baro Corrected sources.

Angle of Attack sources: AaADDU.1 (from the Air Data Computer) was consistently about 0.5 degrees higher than either AA.1 or AA.2, or AaADDU.2. Spike in AA.1 just under 10 seconds prior to takeoff at 7:31:23z (from -186.7 to +20.4 degrees). Similar AA.1 spike just after landing at 15:20:46z (from -150.0 to +81.7 degrees).

GDIFF check: Mean of 1293 meters of drift for INE #1 (with respect to the Novatel)

2014 meters for INE#2. In creation of higher MET parameters, INE #1 was selected.

The Left Dew Point sensor (TD.1) again read far too high through the entire flight (well above ambient temperature). TD.2 did a reasonably good job (staying well below ambient temperature) but is still too high during the cruise portion of flight at or above 41,000 feet (where sondes suggest an RH near 5% and the dew points generated by TD.2 yield an RH around 25%).

For a large portion of cruise, the inertial sources of absolute altitude ran about 400 meters higher than a cluster of the two Honeywell systems and the Novatel. At some portions of cruise, the inertials flipped from 400 meters higher to 150 meters lower than the others.

When plotting PS.1 and PS.2 with PSF.1 and PSF.2 at takeoff, the same small peak in PS.1 and PS.2 appears just after takeoff as seen on all previous WSR-11 missions using AAMPS. On landing approach, there is premature loss of PS.1 and PS.2 correction (associated with the dynamic pressure correction threshold setting). At cruise altitudes, PS.2 (the default) runs about 2 mb higher than PS.1.

Vertical Winds during the cruise portion of flight showed a small low bias (with a mean UZW.1 of -0.13 m/s). As expected, UWZ.1 shows a significant high bias during descent prior to landing.

All other flight level instruments worked optimally during the flight.

- 18 drop points assigned by NCEP

- 20 AVAPS I dropsondes deployed

- Of those 20 drops, 18 were good enough to create a WMO message and were transmitted for ingestion into the 21/12z models. However, two of the good drops were transmitted with no heights: Drop Pt 1 at 08:07:02z failed at 981.5 mb and Drop Pt 7 at 09:47:46z failed at 949.6 mb. There were two Fast Falls (in both cases the backup was good): Pt 12 at 11:17:48z and Pt 13 at 11:39:17z.

- Of the 18 good sondes, 16 coded surface winds.

Flight Directors: Richard Henning and Jessica Williams (813) 828-3310 ext. 3086