NOAA G-IV N49RF ERROR SUMMARY

1

WINTER STORMS 2011

25 Jan 2011 Modified TRACK73 (10WSW)

NCEP (PWT) RJTY RJTY

Flight ID: 20110125N1

Sensor or system						Number or Name

Accelerometer						AccZI.1

Altitude							AltGPS.3

Attack Angle						AA.2

Dew Point Probe						TD.2

Dynamic Pressure						PQF.2

Geopotential Altitude					AltGPS.3

Inertial Selected						INE1

Static Pressure						PSF.2

Slip Angle							SA.1

Temperature Probe				                     TTM.3

True Airspeed						TASF.3

Constants File                                                           49cal102

Flight Directory						acdata/2011/MET/20110125N1

Local Met Data			Takeoff (0757z)		Landing (1549z)

Aircraft Static Pressure		  1001.6 mb			     1004.4 mb

Tower Pressure (corrected)	  1002.9 mb			     1004.3 mb

__________________________________________________________________

Notes:

The only data gaps were in AltRa.1 from the APN-232 radar altimeter:

8:15:20-8:16:05z (multiple gaps during this interval)

15:26:50z

The ALTPA.1 (NACA Pressure Altitude) ran from about 55 meters lower than a tight clustering of Air Data Computer and ADC Baro Corrected sources at cruise early (41K) to 100 meters lower at the highest cruise altitude (45K).

Angle of Attack sources: AaADDU.1 (from the Air Data Computer) was consistently about 0.5 degrees higher than either AA.1 or AA.2, or AaADDU.2. Spike in AA.1 prior to takeoff at 7:56:32z (from -27.2 to +381.7 degrees). Similar AA.1 spike after landing from 15:49:25 - 15:49:26 (-143 to + 201.7 degrees).

GDIFF check: Mean of 1853 meters of drift for INE #1 (with respect to the Novatel)

2264 meters for INE#2. In creation of higher MET parameters, INE #1 selected.

The Left Dew Point sensor (TD.1) again read far too high through the entire flight (well above ambient temperature). TD.2 did a reasonably good job (staying well below ambient temperature) but is still too high during the cruise portion of flight at or above 41,000 feet (where sondes suggest an RH near 5% and the dew points generated by TD.2 yield an RH around 25%).

Vertical Winds looked OK during the cruise portion of flight with a small low bias (a UZW.1 mean of -0.09) but then, as expected, show a significant high bias during descent prior to landing.

All other flight level instruments worked optimally during the flight.

- 19 drop points assigned by NCEP

- 21 AVAPS II dropsondes deployed

- Of those 21 drops, 19 were good. There were two bad sondes (the backup was good at both points): Drop Pt 6 at 10:00:32z had no winds and Drop Pt 10 at 11:13:38z was a Fast Fall.

- Of the 19 good sondes, all 19 coded surface winds.

Flight Directors: Richard Henning and Jack Parrish (813) 828-3310 ext. 3086