NOAA G-IV N49RF ERROR SUMMARY

1

WINTER STORMS 2011

21 Feb 2011 Track 76 (28WSW)

NCEP (PWT) RJTYRJTY

Flight ID: 20110221N1

Sensor or system						Number or Name

Accelerometer						AccZI.1

Altitude							AltGPS.3

Attack Angle						AA.2

Dew Point Probe						TD.2

Dynamic Pressure						PQF.2

Geopotential Altitude					AltGPS.3

Inertial Selected						INE2

Static Pressure						PSF.2

Slip Angle							SA.1

Temperature Probe				                     TTM.1

True Airspeed						TASF.3

Constants File                                                           49cal102

Flight Directory						acdata/2011/MET/20110221N1

Local Met Data			Takeoff (0731z)		Landing (1455z)

Aircraft Static Pressure		  1005.7 mb			     1007.1 mb

Tower Pressure (corrected)	  1008.3 mb			     1009.3 mb

__________________________________________________________________

Notes:

The only data gaps were intervals in AltRa.1 from the APN-232 radar altimeter:

07:45:25-07:46:17z (multiple gaps during this interval during climbout)

08:38:21z

11:59:44z

12:02:20z

In creation of higher MET parameters, PDAlphaF.2 was selected (under Attack Pressure Probe Selection choices) to utilize AA.2.

GDIFF check: Mean of 2896 meters of drift for INE #1 (with respect to the Novatel)

1701 meters for INE#2. In creation of higher MET parameters, INE #2 selected.

The Left Dew Point sensor (TD.1) again read far too high through the entire flight. TD.2 did a reasonably good job but is also too high during the cruise portion of flight at or above 41,000 feet (where sondes suggest an RH near 5% and the dew points generated by TD.2 yield an RH generally between 25 and 35 %).

TTM.3 displayed an oscillation of around 0.4C (with the period of oscillation varying between 3 minutes and 13 minutes) during cruise above 41K. TTM.1 is the default.

Vertical Winds showed a high bias during the cruise portion of flight (with a mean UWZ.1 of +0.30 for the interval from 08:22-13:47z).

All other flight level instruments worked optimally during the flight.

- 18 drop points assigned by NCEP

- 20 AVAPS II dropsondes deployed, (1228z at Drop Pt #14 and 1252z at Drop Pt #15 were fast falls, backups were both good)

- 18 drops good and transmitted for ingestion into 21/1200z model runs

Flight Director: Jessica Williams (813) 828-3310 ext. 3140 and LCDR Nancy Ash