NOAA G-IV N49RF ERROR SUMMARY

1

WINTER STORMS 2011

26 Mar 2011 Reposition

NCEP (PWT) RJTYPHNL

Flight ID: 20110226N1

Sensor or system						Number or Name

Accelerometer						AccZI.1

Altitude							AltGPS.3

Attack Angle						AA.2

Dew Point Probe						TD.2

Dynamic Pressure						PQF.2

Geopotential Altitude					AltGPS.3

Inertial Selected						INE1

Static Pressure						PSF.2

Slip Angle							SA.2

Temperature Probe				                     TTM.1

True Airspeed						TASF.3

Constants File                                                           49cal102

Flight Directory						acdata/2011/MET/20110226N1

Local Met Data			Takeoff (0251z)		Landing (0931z)

Aircraft Static Pressure		  1012.0 mb			     1017.3 mb

Tower Pressure (corrected)	  1014.1 mb			     1017.5 mb

__________________________________________________________________

Notes:

The only data gaps were intervals in AltRa.1 from the APN-232 radar altimeter:

03:04:26-03:04:51z (multiple gaps during this interval during climbout)

08:01:22, 09:05:57, 09:09:08, and 09:09:22z

In creation of higher MET parameters, PDAlphaF.2 was selected (under Attack Pressure Probe Selection choices) to utilize AA.2.

GDIFF check: Mean of 1405 meters of drift for INE #1 (with respect to the Novatel)

2576 meters for INE#2. In creation of higher MET parameters, INE #1 selected.

The Left Dew Point sensor (TD.1) again read far too high through the entire flight. TD.2 did a reasonably good job but is also too high during the cruise portion of flight at or above 41,000 feet (where the dew points generated by TD.2 yield an RH between 25 and 40 %).

TTM.3 displayed an oscillation of around 0.4C (with a period varying for the first several hours of the mission then settling down to a steady 4 minutes) during cruise above 41K. TTM.1 is the default.

Vertical Winds showed a small high bias during the cruise portion of flight (with a mean UWZ.1 of +0.09 for the four hour interval from 04:00-08:07z).

All other flight level instruments worked optimally during the flight.

There were no dropsondes released on this flight (ferry flight).

Flight Director: Jessica Williams (813) 828-3310 ext. 3140 and LCDR Nancy Ash