NOAA G-IV N49RF ERROR SUMMARY

1

WINTER STORMS 2011

3 Mar 2011 Modified TRACK42 (35WSC)

NCEP (PWT) PHNL PHNL

Flight ID: 20110303N1

Sensor or system						Number or Name

Accelerometer						AccZI.1

Altitude							AltGPS.3

Attack Angle						AA.2

Dew Point Probe						TD.2

Dynamic Pressure						PQF.2

Geopotential Altitude					AltGPS.3

Inertial Selected						INE2

Static Pressure						PSF.2

Slip Angle							SA.2

Temperature Probe				                     TTM.1

True Airspeed						TASF.3

Constants File                                                           49cal102

Flight Directory						acdata/2011/MET/20110303N1

Local Met Data			Takeoff (1949z)		Landing (0351z)

Aircraft Static Pressure		  1013.1 mb			     1009.9 mb

Tower Pressure (corrected)	  1014.0 mb			     1011.0 mb

__________________________________________________________________

Notes:

The only data gaps were three intervals in AltRa.1 from the APN-232 radar altimeter:

20:07:58-20:08:37z (multiple gaps during this interval during climbout)

02:04:22z and 03:29:36z

In creation of higher MET parameters, PDAlphaF.2 was selected (under Attack Pressure Probe Selection choices) to utilize AA.2.

GDIFF check: Mean of 2480 meters of drift for INE #1 (with respect to the Novatel)

2130 meters for INE#2. In creation of higher MET parameters, INE #2 selected.

The Left Dew Point sensor (TD.1) again read far too high through the entire flight. TD.2 did a reasonably good job but is also too high during the cruise portion of flight at or above 41,000 feet (where sondes suggest an RH near 5% and the dew points generated by TD.2 yield an RH between 30 and 40 %).

TTM.3 displayed an oscillation of around 0.4C (with a period of approx 4.5 minutes) during cruise above 41K. TTM.1 is the default.

Vertical Winds showed a larger bias than normal during the cruise portion of flight (with a mean UWZ.1 of +0.29 for the six hour interval from 20:44-02:52z).

All other flight level instruments worked optimally during the flight.

- 43 drop points assigned by NCEP

- 44 AVAPS II dropsondes deployed, 1 Fast Fall (2013z Drop Pt #2) Backup was good

- 43 drops good and transmitted for ingestion into 04/0000z model runs

Flight Director: Jessica Williams (813) 828-3310 ext. 3140 and LCDR Nancy Ash