Flight ID: 20110527H1

Sensor or system						Variable Name

Inertial								ine1

Accelerometer							acci1

Temperature Probe						tt1

Dew Point Probe						tdm1

Altitude (for vertical wind)					alti1 (RINU GPS aided inertial alt)

Static Pressure							psfX

Dynamic Pressure						pqf2

Local Met Data:			Takeoff (1116Z)		Landing (1453Z)

Aircraft Static Pressure		 1013.9 mb			1014.9 mb

Tower Pressure			 1013.8 mb			1015.4 mb

Notes

Static Pressure from the fuselage, psf, had stair steps of 1-2mb from 1149Z-1150Z and 1206Z-1208Z.  Dynamic Pressure, pqf1, spikes slightly during these same time periods.  Therefore, pqf2 was selected as default and psf was fixed during the time intervals it spiked using the following substitution: Psf = psw + 3.94mb offset.

All 3 dewpointers had issues.  Tdm3 responds well with the most accurate data, but contains erroneous data spikes to -60C for at least half of the flight.  Tdm2 is way too cold the entire flight averaging at -35C.  Tdm1 is the best, however the dewpoint values at 10k and 12kft are too warm (by 5-7C) and at 5k, 2k and 500ft too cold (also by about 5C) all compared to the sondes dropped in the same area and radiosonde from 12Z tampa sounding. Tdm1 was used as default for this flight.

Novatel GPS, altnvl, contained a spike down 10m at 115813Z.

Gsui and gsvi drift apart from 0.05m/s to 0.6m/s during a roll from 132653Z-132709Z.

AOC and PRO SFMR data was not being displayed on HDOBS.  In realtime, the data appeared higher than observed wind speeds.

All other instruments appeared to work optimally.

6 Dropsondes were released, 3 new ones from NCAR and 3 old (control sondes).  The 3 new sondes were fast falls, the 3 old ones were good.

1 Ch12 AXBT was launched and was good.