1

N42RF ERROR SUMMARY

H. IRENE 0800Z MISSION

OVER THE SE BAHAMAS

Flight ID: 20110824H1

Sensor or system						Number or Name

INE	(for wind derivation)				INE1

Accelerometer						ACCI1

Temperature Probe				          TTM1

Dew Point Probe						TDM2

Static Pressure						PSF

Dynamic Pressure						PQF1

Vert. Wind						          ALTI1

Constants File                                                          /acdata/adc/42_11v5.adc

Project Directory				/acdata/2011/MET/20110824H1

Special note: There was a major problem with the attack angle output during this mission that results in erroneous vertical winds for the entire mission. As of this writing we do not have reason as to why the derived attack angle output was erroneous.

Notes:

There were six data gaps: 092401Z – 092413Z, 103931Z – 103934Z,

113901Z – 113902Z, 115231Z – 115232Z, 123301Z - 123302Z and 124756Z.

RINU-1550 GPS altitude (ALTI1) experienced two separate dropouts at 103935Z and 124757Z. The erroneous data at these separate occurrences was replaced with Novatel GPS (ALTNVL) altitude values.

Other measured parameter output was affected at 103935Z. Ground speed in the

x and y direction (GSUI1 and GSVI1, respectively) and  latitude (LATI1), longitude (LONI1). The erroneous output for each parameter was replaced in the following manner,

GSUI1 = GSUI2 – 0.40

GSVI1 = GSVI2 – 3.20

LATI1 = LATNVL

LONI1 = LONNVL, where NVL refers to Novatel output

Other measured parameter output was affected at 124757Z. Ground speed in the x-direction (GSUI1),  latitude (LATI1), longitude (LONI1), and heading (HDGI1). The erroneous output for each parameter was replaced in the following manner,

GSUI1 = GSUI2 + 1.0

HDGI1 = HDGI2 – 0.60

LATI1 = LATNVL

LONI1 = LONNVL, where NVL refers to Novatel output

Between 091635Z – 091643Z both measured differential attack pressure (PDAF)

and dynamic attack pressure (PQAF) generated erroneous data. For both parameters the erroneous data was removed and replaced using statistical techniques. Radome differential attack pressure (PDAR) was used for modifying PDAF output, and measured fuselage dynamic pressure (PQF1) was used for modifying PQAF output.

Dewpoint sensor 2 (TDM2 [EdgeTech]) generated erroneous data between

141748Z – 141800Z. The erroneous data was replaced with dewpoint sensor 1 (TDM1 [ Buck]) output via direct substitution: TDM2 = TDM1.

Dewpoint sensor 2 (TDM2 [EdgeTech]) generated erroneous data between

160914Z – 161734Z. The erroneous data was replaced with dewpoint sensor  3 (TDM3 [Maycomm TDL]) output via the following equation:

TDM2 = TDM3 + 4.0

During the flight there were instances where dewpoint temperature values exceeded derived ambient temperature values resulting in humidity values above 100%. These situations occurred during heavy precipitation events.

Measured fuselage static pressures, PSF, generated erroneous data between

193500Z – 193700Z.  The erroneous data was removed and replace using statistical techniques with wingtip static pressure (PSW) as a reference.

All other instruments worked optimally during the flight.

There were twenty-eight (28) GPS dropsondes deployed...25 good 3 bad.

SPECIAL NOTE!!!    The variable names dpj_wgs, dpj_was and dpj_wz in the netCDF file represent vertical ground speeds , vertical air speeds and vertical wind speeds, respectively, computed using Dave Jorgensen’s vertical wind algorithm.  It is recommended that these values be used for vertical wind analysis.

Takeoff(0814Z)	Landing(1621Z)

Aircraft Static Pressure                 1011.4mb               1013.2mb

Corrected Tower Pressure             1011.4mb	              1013.0mb

Flight Director:	A. Barry Damiano       (813) 828-3310 ext. 3073