Flight ID: 20110908H1

Sensor or system						Variable Name

Inertial	(for wind derivation)					ine2

Accelerometer							acci2

Temperature Probe						tt1

Dew Point Probe						tdm2

Altitude (for vertical wind)					alti1X

Static Pressure							psfX

Dynamic Pressure						pqf2

Constants file							/acdata/adc/42_11v5.adc

Flight Directory						/acdata/2011/MET/20110908H1

Local Met Data:			Takeoff (1248Z)		Landing (2031Z)

Aircraft Static Pressure		 1013.1 mb			1012.1  mb

Tower Pressure			1013.4  mb			1012.1  mb

Notes

There was a data gap from 132358Z – 132641Z.  Directly after this gap, alti1 exhibited erroneous data for several seconds before it came back online.  Due to this alti1 was replaced with the novatel GPS altitude data, altnvl, for the following time:  Alti1 = Altnvl from 132626Z-132724Z.

All ine1 and acci1 values spiked after this gap as well, so ine2 and acci2 were used as default for calculating wind derivations.

Fuselage static pressure, psf, exhibited 3 downward spikes of 10mb, each spike lasting 2-3 seconds.  The following techniques were used to replace each of the 3 spikes using wingtip static pressure:

171210Z – 171213Z :  psf = psw + 3.27

181103Z – 181106Z :  psf spike removed and replaced with statistical techniques

185847Z – 185850Z :  psf = psw + 3.32

Dynamic pressure pqf1 contained a minor offset from the other dynamic pressure sources from 155518Z – 155528Z, because of this pqf2 was selected as default.

There were periods where the dewpoint temperature exceeded the ambient temperature resulting in RH values > 100%. This was likely due to heavy rain, a wet-bulb effect on the total temperature sensor, and/or an artificial warming of the dewpoint sensor as it tried to burn off excess moisture. Corrections were not made to the data.

AXBT channels 1 and 2 worked properly, where channel 1 had slightly cooler responses for the first 2 hours.

Psf and psw were within 0.5mb of station pressure on takeoff and landing.

All other instruments appeared to work optimally.

A total of 70 instruments were dropped from the aircraft internally:

13 Dropsondes, 10 were good, 3 were fast falls.

19 AXBT;s

24 AXCP’s.

14 AXCTD’s