Flight ID: 20120719N1

Sensor or system						Number or Name

Accelerometer							AccZI.1

Altitude 							AltGPS.3

Attack Angle		PDAlpha.2

Dew Point Probe						TDM.1

Dynamic Pressure						PQM.2

Geopotential Altitude						ALTGA.1 (AltGPS.3)

Inertial Selected						INE.1

Static Pressure							PSM.2

Slip Angle		PDBeta.2

Temperature Probe						TTM.1

Constants File		49cal102

Flight Directory		acdata/MET/2012/20120719N1

Local Met Data:			Takeoff (1519Z)		Landing (1639Z)

Aircraft Static Pressure		1013.2mb			1013.1mb

Tower Pressure (corrected)		1014.1mb			1014.4mb

Notes:

TTM.3 is not in agreement with TTM.1 and TTM.4 and is also not trending with either of the other instruments.  TTM.1 is chosen for the reference.

TASref is using TasADDU.1 in the configuration file.  This is causing gaps throughout the whole flight.  TASref is not used in the calculation of any other parameters.

PQM.1 and PQM.2 are not in good agreement.  It appears that PQM.1 is the faulty instrument when looking at the other dynamic parameters.  PQM.2 should be used.

There are gaps in the AltRa.1 data.  Most noticeably in the multi-hz file (indicated by .fst).  The gaps occur regularly at .465 second intervals along with another gap that is floating at different times.  AltRa.1 was very noisy at FL300.

TDM.1 and TDM.2 are both not rated for use under -50 deg C.  Cannot be considered reliable for dewpoints colder than -50C.

AltGPS.3 (Novatel) was not online until 10:50:31z.  It took some troubleshooting by the techs to find a loose cable connection but from there on it worked fine.