NOAA WP-3D N42RF ERROR SUMMARY

1

OCEAN WINDS WINTER 2013

22 Jan 2013 PROJECT FLIGHT #2

NESDIS (PND) CYHZ  CYHZ

Flight ID: 20130122H1

Sensor or system						Number or Name

Altitude							AltIGPS.1 (RINU)

Accelerometer						AccZfilterI-GPS.1

Dew Point Probe						TDM.2

Dynamic Pressure						PQM.2

Inertial Selected						INE1

Static Pressure						PSM.2

Temperature Probe				                     TTM.1

Constants File                                                           AAMPSConfig/core/n43.xml

Flight Directory						acdata/MET/20130122H1

Local Met Data			Takeoff (1607z)		Landing (1946z)

Aircraft Static Pressure		  985.7 mb			     986.2 mb

Tower Pressure (corrected)	  987.1 mb			     985.7 mb

__________________________________________________________________

Notes:

Science portion of mission from 15:16z – 18:52z flown at 7000 ft radar altitude. There was a 40 second gap in all AAMPS data from 16:55:56z to 16:56:36z.

Maycomm (TDM.3) dew point sensor was essentially unusable. TDM.2 (Edgetech), the reference, as well as TDM.1 (Buck) both performed very well throughout the mission.

All other flight level instruments worked optimally during the flight.

Vertical Winds during science portion showed a mean UWZ of -0.02 m/s. The variable names GSZ_DPJ, ASZ_DPJ and WSZ_DPJ in the netCDF file represent vertical ground speeds vertical air speeds and vertical wind speeds, respectively, computed using Dave Jorgensen’s vertical wind algorithm. It is recommended that these values be used for vertical wind analysis.

- 4 dropsondes deployed, all 4 good, none transmitted

- 2 internally loaded AXBT deployed both good

Flight Directors: Richard Henning (813) 828-3310 ext. 3086 and A Barry Damiano