NOAA WP-3D N42RF ERROR SUMMARY

1

OCEAN WINDS WINTER 2013

23 Jan 2013 PROJECT FLIGHT #3

NESDIS (PND) CYHZ  CYQX (refuel)

Flight ID: 20130123H1

Sensor or system						Number or Name

Altitude							AltIGPS.1 (RINU)

Accelerometer						AccZfilterI-GPS.1

Dew Point Probe						TDM.2X

Dynamic Pressure						PQM.2

Inertial Selected						INE1

Static Pressure						PSM.2

Temperature Probe				                     TTM.1

Constants File                                                           AAMPSConfig/core/n43.xml

Flight Directory						acdata/MET/20130123H1

Local Met Data			Takeoff (1245z)		Landing (2025z)

Aircraft Static Pressure		  992.3 mb			     980.0 mb

Tower Pressure (corrected)	  992.8 mb			     978.3 mb

__________________________________________________________________

Notes:

Science portion of mission from 15:16z – 18:52z flown at 7000 to 7500 ft radar altitude.

Maycomm (TDM.3) dew point sensor was essentially unusable. TDM.2 (Edgetech), the reference, ran away during the high altitude transit prior to the science portion. From 13:15z to 14:27z values for TDM.1 (Buck) were substituted.

All other flight level instruments worked optimally during the flight.

Vertical Winds during the science portion showed a mean UWZ of -0.14 m/s. The variable names GSZ_DPJ, ASZ_DPJ and WSZ_DPJ in the netCDF file represent vertical ground speeds vertical air speeds and vertical wind speeds, respectively, computed using Dave Jorgensen’s vertical wind algorithm. It is recommended that these values be used for vertical wind analysis.

- 22 dropsondes deployed, all 22 good, none transmitted

- 1 internally loaded AXBT deployed was good

Flight Directors: Richard Henning (813) 828-3310 ext. 3086 and A Barry Damiano